



	
[image: image1.png]



	Medical Devices Control Division

Ministry of Health Malaysia
Level 5, Plot 3C4, 

No. 26, Jalan Persiaran Perdana,

Precint 3, 62675 Putrajaya
Malaysia
Tel: 603-8885 0778 / 603-8885 0773
Fax: 603-8885 0758



Medical Device / Equipment VOLUNTARY RECALL UPDATES
Date Issued : 27th OCTOBER 2010                         Ref:MDB/R/2010/011
	IMMEDIATE ACTION
	

	ACTION
	√

	UPDATE
	√

	INFORMATION REQUEST
	


*The dissemination of this information is a follow up from our Recall Notice issued on 18th October 2010
	PRODUCT
	Various Medical Devices Bearing COOK brand name.
Please see attachments for details.



	CLASS
	Class I

	USE
	Various Usage

	SOURCE OF MEDICAL DEVICE RECALL / ALERT
	USFDA website:
http://www.fda.gov/MedicalDevices/Safety/RecallsCorrectionsRemovals/ListofRecalls/ucm221229.htm
Recall Initiated: 21st April 2010



	RECALLING / ALERTING FIRM 
	MALAYSIA REPRESENTATIVE

Director (Mr. Lee Hong Guan)

Cook Asia (Malaysia) Sdn. Bhd.,

48-2A, Block E,

Jalan PJU 1/3B,

47301 Petaling Jaya, Selangor

MALAYSIA

T- 603 7880 4485

F- 603 7880 1910



	REASON FOR RECALL/ALERT
	The set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC Tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
*These sets and trays were distributed from February 18, 2009 through April 27, 2010.



	SCENARIO IN MALAYSIA
	From e-mail communication with Cook Asia (Malaysia) Sdn. Bhd. 
The affected products were returned by the customers during the exercise done in the month of June 2010
Total affected number = 14 sets

All affected number was sent back to the USA on 15th June 2010.



	ACTION 
	*any individual possessing the abovementioned device are advised to respond to this communication immediately.


	RECOMMENDATION
	Users of the abovementioned device should contact the distributors/supplier of this device/product and inform the Medical Devices Control Division, Ministry of Health providing the following information:-


a. Contact person and contact number


b. Numbers of units available
            c. Name of the supplier(s)



	CONTACT/ENQUIRIES
	Customers may contact Cook Medical Customer Relations Department at 1-800-457-4500 or 1-812-339-2235, or by email at CustomerRelationsNA@CookMedical.com.


	REFERENCES
	http://www.fda.gov/Safety/Recalls/ucm219722.htm


PAGE  
THE INFORMATION PUBLISHED IN THIS WEBSITE IS NOT VALID FOR ANY MARKETING PURPOSE AND STRATEGY. 
Page 1

_1011766759.doc
[image: image1.png]






