

	
[image: image1.png]



	Medical Devices Control Division

Ministry of Health Malaysia
Level 5, Plot 3C4, 

No. 26, Jalan Persiaran Perdana,

Precint 3, 62675 Putrajaya
Malaysia
Tel:  603-8885 0778 
Fax: 603-8885 0758



Medical Device / Equipment VOLUNTARY RECALL 
Date Issued : 26th JANUARY 2012                         Ref:MDB/R/2012/001
	IMMEDIATE ACTION
	√

	ACTION
	

	UPDATE
	

	INFORMATION REQUEST
	√


*The dissemination of this information is based on the initial information gathered via USFDA website
	PRODUCT
	RESPIRONICS TRILOGY 100 VENTILATORS
*please refer to the link provided for affected serial numbers



	RECALL CLASS
	Class 1

	DEVICE CLASS
	n/a

	USE
	Ventilators

	SOURCE OF MEDICAL DEVICE RECALL / ALERT
	USFDA website:

http://www.fda.gov/MedicalDevices/Safety/RecallsCorrectionsRemovals/ListofRecalls/ucm287528.htm 

	RECALLING / ALERTING FIRM 
	Respironics Inc.

	REASON FOR RECALL/ALERT
	Trilogy 100 ventilators are being recalled due to a manufacturing issue. This issue can cause part of the blower that circulates air and other gases through the ventilator to move out of position and cause the device to alarm and stop delivering therapy to the patient. In the event of a blower failure, the ventilator will stop delivering therapy and a high priority alarm will sound to alert the caregiver to immediately respond. Failure to respond could result in the potential for harm or death of a ventilator-dependent patient.

	SCENARIO IN MALAYSIA
	NO INFORMATION AVAILABLE AT THE MOMENT

	ACTION 
	*any healthcare facilities possessing the abovementioned device are advised to respond to this communication immediately.



	RECOMMENDATION
	Users of the abovementioned device should contact the distributors/supplier of this device/product and inform the Medical Devices Control Division, Ministry of Health Malaysia providing the following information:-


a. Name of healthcare centre/hospital/clinic


b. Contact person and contact number


c. Numbers of units available

            d. Name of the supplier(s)



	CONTACT/ENQUIRIES
	N/A

	REFERENCES
	-Please Refer Attachments-




PAGE  
 INFORMATION PUBLISHED IN THIS WEBSITE IS NOT VALID FOR ANY MARKETING PURPOSE AND STRATEGY. 
Page 1

_1011766759.doc
[image: image1.png]






